Analytical Method Development and Validation for the Determination of Ethylene Glycol and Diethylene Glycol in Herbal Cough Syrup by Gas Chromatography

Ankit Sharma, Subrahmanya Kumar Kukkupuni

Centre for Ayurveda Biology and Holistic Nutrition,

The University of Trans-Disciplinary Health Sciences and Technology, Bengaluru, India

Corresponding Author Email: ankitissyogi@gmail.com, s.kumar@tdu.edu.in,

Abstract

Ethylene Glycol (EG) and Diethylene Glycol (DEG) are potential contaminants in liquid oral formulations, posing severe toxicological risks. Regulatory agencies mandate stringent controls on their permissible limits. This study outlines the analytical method verification of a gas chromatography (GC) procedure for the quantitative determination of EG and DEG in a polyherbal cough syrup. The method employs a DB-624 capillary column and flame ionization detection (FID), using ethanol as the diluent. System suitability, precision, resolution, and tailing factor criteria were established and verified. The method was confirmed to be accurate, precise, and suitable for routine use in quality control.

Keywords: Ethylene Glycol, Diethylene Glycol, Gas Chromatography, Method Verification, Cough Syrup, System Suitability

1. Introduction

Ethylene Glycol (EG) and Diethylene Glycol (DEG) are used as industrial solvents and antifreeze agents. However, their presence in pharmaceutical formulations, whether due to contamination or substitution, can lead to serious toxicity, including renal failure and death [1]. Several global incidents have drawn attention to the need for accurate and validated methods to detect these compounds in pharmaceutical products [2].

Naturals Cough Syrup is a polyherbal formulation, and it is essential to ensure that it is free from toxic contaminants like EG and DEG. This study describes the verification of an analytical method employing Gas Chromatography (GC) for the quantitative estimation of EG and DEG, as per ICH Q2(R1) and applicable regulatory guidelines.

2. Materials and Methods

2.1 Chemicals and Reagents

- Ethanol (HPLC Grade, Merck)

- Ethylene Glycol (Reference Standard)

- Diethylene Glycol (Reference Standard)

- Propylene Glycol (Internal Standard)

2.2 Instruments and Equipment

- Gas Chromatograph: Agilent 7820 / 7860 with FID

- Column: DB-624, 30 m x 0.25 mm ID x 1.4 µm film thickness

- Analytical Balance: Shimadzu (Actzec)

- Micropipettes: Thermo Fisher

- Sonicator: Labman

2.3 Sample Preparation 10 g of Naturals Cough Syrup was transferred into a 100 mL volumetric flask, sonicated for 10 minutes, and volume made up with ethanol. The solution was filtered through a 0.45 µm syringe filter before injection.

2.4 Standard and System Suitability Solutions

- Standard Stock: 100 mg each of EG and DEG was dissolved in ethanol up to 100 mL.

- Working Standard: 5 mL of the stock was diluted to 50 mL (100 ppm).

- System Suitability: 5 mL of EG/DEG and 5 mL of Propylene Glycol stock were diluted to 50 mL with ethanol.

2.5 Chromatographic Conditions

| Parameter | Value |
| --- | --- |
| Column | DB-624 |
| Injector Temp | 230°C |
| Detector Temp (FID) | 250°C |
| Injection Volume | 1 µL |
| Split Ratio | 1:20 |
| Carrier Flow | 2 mL/min |

3. Analytical method validation

3.1 Specificity

3.2 Precision

3.2.1 System Precision

3.2.2 Method Precision

3.3 Limit of Detection and Limit of Quantitation

3.4 Accuracy

3.5 Solution Stability

3.1 Specificity

The specificity of the developed Gas Chromatographic (GC) method was demonstrated by evaluating potential interferences from various excipients, herbal actives, and formulation components. The test solutions included blank, standard (EG and DEG), placebo, spiked sample, and individual solutions of herbal and excipient constituents to ensure there was no co-elution or interference with the analyte peaks of EG and DEG.

#### Preparation of Solutions for Specificity Study

1. Propylene Glycol Stock Solution:
200.5 mg of Propylene Glycol was accurately weighed into a 100 mL volumetric flask. Approximately 70 mL of ethanol was added, and the mixture was shaken until completely dissolved. The final volume was made up with ethanol.

2. EG and DEG Stock Solution A:
102.8 mg of Ethylene Glycol and 101.5 mg of Diethylene Glycol were transferred into a 100 mL volumetric flask. After the addition of 70 mL of ethanol and thorough mixing, the final volume was adjusted with ethanol.

3. Standard Solution A:
5 mL of EG & DEG Stock Solution A was diluted to 50 mL with ethanol and mixed well before injection.

4. EG and DEG Stock Solution B:
101.4 mg of EG and 100.4 mg of DEG were similarly prepared in a 100 mL flask and diluted with ethanol as above.

5. Standard Solution B:
5 mL of Stock Solution B was diluted to 50 mL with ethanol, mixed, and injected.

6. System Suitability Solution:
5 mL each of Propylene Glycol and EG & DEG stock solutions were mixed in a 50 mL volumetric flask, diluted with ethanol, and injected to check system suitability parameters.

#### Placebo and Sample Solutions

7. Placebo Solution:
10.0219 g of placebo was dissolved in ethanol using sonication for 10 minutes, made up to 100 mL, filtered (0.45 µm syringe filter), and injected.

8. As Such Sample Solutions (1 & 2):
10.0425 g of the syrup sample was weighed and treated identically as a placebo for both preparations. Solutions were filtered and injected.

9. Spiked Sample Solution:
10.0368 g of sample was spiked with 10 mL of EG & DEG standard stock solution, sonicated, volume made up with ethanol, filtered, and injected.

10. Unfiltered Sample Solutions (1 & 2):
Samples of 10.0516 g and 10.0159 g were prepared without filtration, diluted in ethanol, and directly injected after mixing.

#### Identification Solutions of Herbal Ingredients and Excipients

Each of the following components was weighed in the range of ~100 mg, dissolved in 70 mL of ethanol, and the volume was made up to 100 mL with ethanol. From this, 5 mL was further diluted to 50 mL with ethanol for injection. The purpose was to assess their chromatographic behaviour and potential interference.

| Component | Quantity (mg) |
| --- | --- |
| Ethylene Glycol | 100.7 |
| Diethylene Glycol | 100.1 |
| Madhu (Honey) | 103.1 |
| Yashtimadhu (Glycyrrhiza glabra) | 102.01 |
| Tulsi (Ocimum sanctum) | 101.08 |
| Vasaka (Adhatoda vasica) | 100.41 |
| Talishpatra (Taxus baccata) | 102.42 |
| Sunthi (Zingiber officinale) | 100.85 |
| Haldi (Curcuma longa) | 100.93 |
| Banashpa (Viola odorata) | 100.14 |
| Pippali (Piper longum) | 100.27 |
| Shati (Hedychium spicatum) | 102.37 |
| Munkka (Vitis vinifera) | 101.90 |
| Sucrose | 99.57 |
| Sorbitol Solution 70% | 102.09 |
| Liquid Glucose | 101.40 |
| Propylene Glycol | 100.3 |
| Citric Acid Monohydrate | 100.80 |
| Sodium Benzoate | 102.50 |
| Sodium Methyl Paraben | 99.43 |
| Menthol | 102.14 |
| Sodium Propyl Paraben | 102.38 |
| Oleo-resin (Ginger) | 102.6 |
| Camphor | 100.03 |
| Purified Water | 102.9 |

Each solution was carefully prepared by weighing the component, dissolving in ethanol, sonication if needed, final volume makeup, dilution (5 mL into 50 mL ethanol), and analysis through GC.

#### System Suitability

System suitability parameters were evaluated by injecting standard solutions (Standard Solution-A and Standard Solution-B) and analysing critical attributes such as similarity factor, resolution, %RSD for area and retention time, and tailing factors for both Ethylene Glycol and Diethylene Glycol.

Table 2: System Suitability Evaluation

| System Suitability Parameter | Analyte | Observation | Acceptance Criteria |
| --- | --- | --- | --- |
| Similarity Factor (Std A vs. Std B) | Ethylene Glycol | 0.96 | NLT 0.85 and NMT 1.15 |
|  | Diethylene Glycol | 0.99 | NLT 0.85 and NMT 1.15 |
| Resolution between EG and PG peaks | - | 13.09 | NLT 10 |
| % RSD of Peak Area (n=6) | Ethylene Glycol | 8.2 | NMT 15% |
|  | Diethylene Glycol | 6.6 | NMT 15% |
| % RSD of Retention Time (n=6) | Ethylene Glycol | 0.02 | NMT 2% |
|  | Diethylene Glycol | 0.06 | NMT 2% |
| Tailing Factor | Ethylene Glycol | 1.6 | NMT 3.5 |
|  | Diethylene Glycol | 2.0 | NMT 3.5 |

All parameters met the pre-defined acceptance criteria, confirming the system suitability for method validation.

#### Specificity

To confirm the method’s specificity, retention times for Ethylene Glycol and Diethylene Glycol were determined in the presence of matrix and excipients. No interfering peaks were observed at these retention times.

Table 3: Retention Time of Target Analytes

| Peak Name | Retention Time (min) |
| --- | --- |
| Ethylene Glycol | 5.987 |
| Diethylene Glycol | 17.179 |

#### Filter Study

The purpose of the filter study was to confirm the absence of any loss of analyte or interference introduced by filtration. Both filtered and unfiltered sample solutions were analyzed, and no significant difference in response was observed.

Table 4: Filter Study Data

| Injection | Ethylene Glycol Area | Diethylene Glycol Area |
| --- | --- | --- |
| Sample – As Such – 1 (Without Filter) | 0 | 0 |
| Sample – As Such – 2 (Without Filter) | 0 | 0 |
| Sample – As Such – 1 (Filtered) | 0 | 0 |
| Sample – As Such – 2 (Filtered) | 0 | 0 |

Conclusion

- The method passed all system suitability criteria, indicating consistent and reliable chromatographic performance.

- No extraneous peaks were observed in the filtered sample solution, confirming specificity and absence of matrix interference.

- There was no difference in analyte response between filtered and unfiltered sample solutions, establishing filter compatibility.

### 3.2 Precision

#### 3.2.1 System Precision

System precision was evaluated by performing six replicate injections of the standard solution containing Ethylene Glycol (EG) and Diethylene Glycol (DEG). The aim was to demonstrate the reproducibility of the chromatographic system under the defined operating conditions.

Preparation of Solutions

- Propylene Glycol Stock Solution:
Accurately weighed 201.2 mg of Propylene Glycol into a 100 mL volumetric flask. Added 70 mL ethanol, mixed thoroughly, and diluted to volume with ethanol. The solution was shaken well before use.

- EG & DEG Stock Solution-A:
Transferred 100.4 mg of Ethylene Glycol and 102.1 mg of Diethylene Glycol into a 100 mL volumetric flask. Dissolved in 70 mL of ethanol and diluted to volume with ethanol. The solution was shaken well to ensure homogeneity.

- Standard Solution-A:
5 mL of EG & DEG Stock Solution-A was diluted to 50 mL with ethanol, shaken well, and injected.

- EG & DEG Stock Solution-B:
Prepared similarly using 101.1 mg of EG and 101.9 mg of DEG in ethanol to a final volume of 100 mL.

- Standard Solution B:
5 mL of EG & DEG Stock Solution-B was diluted to 50 mL with ethanol and injected.

- System Suitability Solution (EG, DEG & Propylene Glycol):
Combined 5 mL each of Propylene Glycol and EG & DEG stock solutions in a 50 mL volumetric flask. Diluted to volume with ethanol and injected to verify suitability.

#### System Suitability Results

Table 5: System Suitability Parameters

| Parameter | Analyte | Observation | Acceptance Criteria |
| --- | --- | --- | --- |
| Similarity Factor (Standard A vs. B) | Ethylene Glycol | 1.01 | NLT 0.85 and NMT 1.15 |
|  | Diethylene Glycol | 1.02 | NLT 0.85 and NMT 1.15 |
| Resolution between EG and Propylene Glycol peaks | - | 12.96 | NLT 10 |
| % RSD of Area (n=6) | Ethylene Glycol | 2.2 | NMT 15% |
|  | Diethylene Glycol | 4.4 | NMT 15% |
| % RSD of Retention Time (n=6) | Ethylene Glycol | 0.02 | NMT 2% |
|  | Diethylene Glycol | 0.05 | NMT 2% |
| Tailing Factor | Ethylene Glycol | 1.7 | NMT 3.5 |
|  | Diethylene Glycol | 2.3 | NMT 3.5 |

#### System Precision Data

Table 6: Precision – Retention Time and Peak Area (n = 6)

| Injection | Ethylene Glycol |  | Diethylene Glycol |  |
| --- | --- | --- | --- | --- |
|  | Retention Time (min) | Area | Retention Time (min) | Area |
| Standard Solution_01 | 5.988 | 190,875 | 17.174 | 190,241 |
| Standard Solution_02 | 5.990 | 191,365 | 17.192 | 195,484 |
| Standard Solution_03 | 5.989 | 191,505 | 17.192 | 191,174 |
| Standard Solution_04 | 5.988 | 191,949 | 17.173 | 192,114 |
| Standard Solution_05 | 5.988 | 193,698 | 17.174 | 189,143 |
| Standard Solution_06 | 5.989 | 196,172 | 17.189 | 194,394 |
| Average | 5.989 | 192,594 | 17.182 | 192,092 |
| % RSD | 0.01 | 1.0 | 0.06 | 1.3 |

#### Conclusion

- All system suitability parameters complied with pre-defined acceptance criteria, indicating a stable and reliable chromatographic system.

- The % Relative Standard Deviation (RSD) for peak areas of Ethylene Glycol and Diethylene Glycol in six replicate injections was found to be 1.0% and 1.3%, respectively, well within the acceptable limit of 15%.

- The % RSD for retention times was 0.01% for EG and 0.06% for DEG, which is significantly below the 2% acceptance threshold.

- These results confirm the system precision of the method for simultaneous estimation of EG and DEG in the matrix.

### 3.2.2 Method Precision

Method precision, or repeatability, was evaluated by preparing and analyzing six individual spiked sample solutions of Wet Cough Syrup from a single batch. Each sample was spiked with known concentrations of Ethylene Glycol (EG) and Diethylene Glycol (DEG), and the % Relative Standard Deviation (%RSD) for both analytes was calculated to assess the consistency of the method under normal operating conditions.

#### Preparation of Solutions

- Propylene Glycol Stock Solution:
Accurately weighed 201.2 mg of Propylene Glycol into a 100 mL volumetric flask. Added ~70 mL ethanol, mixed well, and made up to volume with ethanol. The solution was shaken thoroughly.

- EG & DEG Stock Solution-A:
Transferred 100.4 mg of Ethylene Glycol and 102.1 mg of Diethylene Glycol into a 100 mL volumetric flask, dissolved in ethanol (70 mL), and made up to volume with ethanol. The solution was shaken well.

- Standard Solution-A:
5 mL of EG & DEG Stock Solution-A was diluted to 50 mL with ethanol, mixed, and injected.

- EG & DEG Stock Solution-B:
Weighed 101.1 mg of EG and 101.9 mg of DEG into a 100 mL flask, prepared as above.

- Standard Solution B:
5 mL of Stock Solution-B was diluted to 50 mL with ethanol and injected.

- System Suitability Solution (EG, DEG & Propylene Glycol):
5 mL each of EG & DEG and Propylene Glycol stock solutions were diluted to 50 mL with ethanol, mixed well, and injected.

#### Preparation of Spiked Sample Solutions

Six individual spiked samples were prepared by weighing approximately 10 g of cough syrup and adding 10 mL of EG & DEG Standard Stock Solution-A, followed by dilution to 100 mL with ethanol.

Table 7: Spiked Sample Preparations

| Sample Name | Weight (g) | Stock Solution Added (mL) | Final Volume (mL) | Remarks |
| --- | --- | --- | --- | --- |
| Spiked Sample Solution_1 | 10.0390 | 10 | 100 | Prepared as per Section 6.7 |
| Spiked Sample Solution_2 | 10.1637 | 10 | 100 |  |
| Spiked Sample Solution_3 | 10.0226 | 10 | 100 |  |
| Spiked Sample Solution_4 | 10.0335 | 10 | 100 |  |
| Spiked Sample Solution_5 | 10.0069 | 10 | 100 |  |
| Spiked Sample Solution_6 | 10.1333 | 10 | 100 |  |

#### Results

##### System Suitability Evaluation

Table 8: System Suitability Parameters

| Parameter | Analyte | Observation | Acceptance Criteria |
| --- | --- | --- | --- |
| Similarity Factor (Standard A vs. B) | Ethylene Glycol | 1.01 | NLT 0.85 and NMT 1.15 |
|  | Diethylene Glycol | 1.02 | NLT 0.85 and NMT 1.15 |
| Resolution between EG and Propylene Glycol peaks | - | 12.96 | NLT 10 |
| % RSD of Area (n=6) | Ethylene Glycol | 1.3 | NMT 15% |
|  | Diethylene Glycol | 5.4 | NMT 15% |
| % RSD of Retention Time (n=6) | Ethylene Glycol | 0.01 | NMT 2% |
|  | Diethylene Glycol | 0.06 | NMT 2% |
| Tailing Factor | Ethylene Glycol | 1.7 | NMT 3.5 |
|  | Diethylene Glycol | 2.3 | NMT 3.5 |

##### Method Precision – Spiked Sample Results

Table 9: Recovery and %RSD of Spiked Samples (n = 6)

| Sample | Ethylene Glycol (% Content) | Recovery (%) | Diethylene Glycol (% Content) | Recovery (%) |
| --- | --- | --- | --- | --- |
| Spiked Sample 1 | 0.103 | 103.0 | 0.113 | 112.5 |
| Spiked Sample 2 | 0.108 | 107.8 | 0.117 | 117.1 |
| Spiked Sample 3 | 0.110 | 109.6 | 0.118 | 118.2 |
| Spiked Sample 4 | 0.109 | 109.1 | 0.119 | 118.6 |
| Spiked Sample 5 | 0.110 | 109.8 | 0.119 | 119.1 |
| Spiked Sample 6 | 0.109 | 109.1 | 0.119 | 119.3 |
| Average | 0.108 | 108.1 | 0.117 | 117.5 |
| % RSD | 2.4 | 2.4 | 2.2 | 2.2 |

#### Conclusion

- All system suitability parameters were within acceptable limits, confirming chromatographic system performance.

- The %RSD of Ethylene Glycol and Diethylene Glycol content across six replicate spiked samples was 2.4% and 2.2%, respectively—both significantly below the acceptable limit of 15%.

- Recovery values for both analytes ranged between 103.0% to 119.3%, indicating the method is accurate for the quantification of EG and DEG at the specified concentration levels.

- These results confirm that the method is precise and reproducible for the analysis of Ethylene Glycol and Diethylene Glycol in Wet Cough Syrup.

### 3.3 Limit of Detection (LOD) and Limit of Quantitation (LOQ)

The Limit of Detection (LOD) and Limit of Quantitation (LOQ) for Ethylene Glycol (EG) and Diethylene Glycol (DEG) were determined to establish the sensitivity of the analytical method. These parameters were calculated based on the residual standard deviation of the response and the slope of the calibration curve generated from standard solutions in the concentration range of 5% to 30% of the standard concentration.

Predicted LOD and LOQ Values

Table 10: Calculated LOD and LOQ

| Analyte | LOD (% w/w) | LOQ (% w/w) |
| --- | --- | --- |
| Ethylene Glycol | 0.001 | 0.003 |
| Diethylene Glycol | 0.001 | 0.003 |

#### Preparation of Solutions

- Propylene Glycol Stock Solution:
201.2 mg of Propylene Glycol was dissolved in ethanol and made up to 100 mL.

- EG & DEG Stock Solution-A:
100.4 mg of EG and 102.1 mg of DEG were dissolved in ethanol and made up to 100 mL.

- Standard Solution-A & B:
5 mL of the stock solution was diluted to 50 mL with ethanol.

- System Suitability Solution:
5 mL each of Propylene Glycol and EG/DEG stock solutions were diluted to 50 mL.

- LOD Solution:
10.0245 g of placebo was spiked with 1.0 mL of Standard Stock Solution-A, sonicated, diluted to 100 mL with ethanol, filtered (0.45 µm), and injected.

- LOQ Solution:
10.0359 g of placebo was spiked with 3.0 mL of Standard Stock Solution-A, processed as above, and injected.

#### Results

##### System Suitability

Table 11: System Suitability Criteria

| Parameter | Analyte | Observation | Acceptance Criteria |
| --- | --- | --- | --- |
| Similarity Factor (Standard A vs. B) | EG / DEG | 1.01 / 1.02 | NLT 0.85 – NMT 1.15 |
| Resolution (EG vs. Propylene Glycol) | - | 12.96 | NLT 10 |
| % RSD of Peak Area (n=6) | EG / DEG | 2.2 / 4.4 | NMT 15% |
| % RSD of Retention Time (n=6) | EG / DEG | 0.02 / 0.05 | NMT 2% |
| Tailing Factor | EG / DEG | 1.7 / 2.3 | NMT 3.5 |

##### Limit of Detection (LOD)

Table 12: LOD S/N Ratios

| Injection | EG S/N Ratio | DEG S/N Ratio |
| --- | --- | --- |
| LOD_1 | 20.07 | 7.05 |
| LOD_2 | 16.55 | 6.68 |
| LOD_3 | 15.55 | 7.40 |

Observation: Signal-to-noise ratios for both analytes exceed 3, confirming detectability at the LOD level.

##### Limit of Quantitation (LOQ)

Table 13: LOQ Precision and S/N Ratios (n = 6)

| Injection | EG RT (min) | EG Area | EG S/N | DEG RT (min) | DEG Area | DEG S/N |
| --- | --- | --- | --- | --- | --- | --- |
| LOQ_1 | 6.004 | 51,538 | 70.76 | 17.193 | 65,948 | 25.28 |
| LOQ_2 | 6.004 | 50,179 | 72.07 | 17.194 | 66,718 | 25.98 |
| LOQ_3 | 6.003 | 52,145 | 70.74 | 17.189 | 66,128 | 25.20 |
| LOQ_4 | 6.002 | 50,413 | 72.44 | 17.191 | 64,528 | 25.39 |
| LOQ_5 | 6.002 | 50,445 | 69.07 | 17.199 | 67,413 | 24.82 |
| LOQ_6 | 6.001 | 52,243 | 73.01 | 17.187 | 64,636 | 25.26 |
| Average | 6.003 | 51,161 | — | 17.181 | 65,895 | — |
| %RSD | 0.02 | 1.8 | — | 0.02 | 1.7 | — |

#### Conclusion

- The system suitability parameters met all acceptance criteria, validating the chromatographic system's performance.

- The S/N ratio for LOD was consistently > 3, and for LOQ, it was > 10, satisfying regulatory thresholds.

- The %RSD of area for EG and DEG in six injections at the LOQ level was 1.8% and 1.7%, respectively, well below the 15% threshold.

- The %RSD of retention time was 0.02% for both EG and DEG, confirming consistent chromatographic behaviour.

- These results affirm the method’s sensitivity and precision at low analyte concentrations.

### 3.4 Accuracy / Recovery

The accuracy of the developed gas chromatographic method was evaluated for the estimation of Ethylene Glycol (EG) and Diethylene Glycol (DEG) in Natural Cough Syrup. Accuracy studies were conducted by recovery experiments at four levels: LOQ, 50%, 100%, and 150% of the specification concentration. Recovery was calculated by comparing the measured concentration of spiked samples with the theoretical concentration.

#### Preparation of Solutions

- Propylene Glycol Stock Solution:
201.2 mg of Propylene Glycol was transferred into a 100 mL volumetric flask, dissolved in 70 mL of ethanol, and the volume was made up with ethanol. The solution was mixed thoroughly.

- EG & DEG Stock Solution-A:
100.4 mg of EG and 102.1 mg of DEG were weighed and similarly diluted to 100 mL with ethanol.

- Standard Solutions A and B:
5 mL of EG & DEG stock solution was diluted to 50 mL with ethanol and mixed well before injection.

- System Suitability Solution:
A mixture of 5 mL of Propylene Glycol stock solution and 5 mL of EG & DEG stock solution was diluted to 50 mL with ethanol and injected.

- Spiking Stock Solution:
107.4 mg of EG and 108.3 mg of DEG were dissolved and diluted to 100 mL with ethanol to prepare the standard stock solution for spiking.

- Sample Solutions:
The placebo matrix was spiked with the standard stock solution at defined volumes to achieve the desired % levels. Sample weights and spiking details are summarised in the respective tables below.

### Results

#### System Suitability

Table 14: System Suitability Parameters

| Parameter | Analyte | Observation | Acceptance Criteria |
| --- | --- | --- | --- |
| Similarity Factor (Standard A vs. B) | EG / DEG | 1.01 / 1.02 | 0.85 – 1.15 |
| Resolution between the EG and Propylene Glycol peaks | - | 12.96 | NLT 10 |
| % RSD of peak areas (n = 6) | EG / DEG | 2.3 / 6.7 | NMT 15% |
| % RSD of retention time (n = 6) | EG / DEG | 0.03 / 0.06 | NMT 2% |
| Tailing Factor | EG / DEG | 1.7 / 2.3 | NMT 3.5 |

#### Accuracy / Recovery of Ethylene Glycol

Table 15: Accuracy of Ethylene Glycol

| Level | Sample No. | Area | Recovery (%) | Mean Recovery (%) | RSD (%) |
| --- | --- | --- | --- | --- | --- |
| As Such | 1 | ND | NA | NA | NA |
|  | 2 | ND | NA |  |  |
|  | 3 | ND | NA |  |  |
| LOQ | 1 | 73517 | 118.9 | 115.1 | 3.0 |
|  | 2 | 69462 | 112.4 |  |  |
|  | 3 | 70405 | 113.9 |  |  |
| 50% | 1 | 108457 | 105.3 | 105.7 | 0.4 |
|  | 2 | 109052 | 108.9 |  |  |
|  | 3 | 109254 | 106.1 |  |  |
| 100% | 1 | 211092 | 102.9 | 103.3 | 0.5 |
|  | 2 | 212776 | 103.3 |  |  |
|  | 3 | 213907 | 103.8 |  |  |
| 150% | 1 | 322894 | 104.5 | 102.5 | 1.7 |
|  | 2 | 314748 | 101.8 |  |  |
|  | 3 | 312999 | 101.3 |  |  |

#### Accuracy / Recovery of Diethylene Glycol

Table 16: Accuracy of Diethylene Glycol

| Level | Sample No. | Area | Recovery (%) | Mean Recovery (%) | RSD (%) |
| --- | --- | --- | --- | --- | --- |
| As Such | 1 | ND | NA | NA | NA |
|  | 2 | ND | NA |  |  |
|  | 3 | ND | NA |  |  |
| LOQ | 1 | 72948 | 119.3 | 118.0 | 1.3 |
|  | 2 | 71138 | 116.4 |  |  |
|  | 3 | 72305 | 118.3 |  |  |
| 50% | 1 | 113281 | 111.2 | 111.3 | 1.1 |
|  | 2 | 114683 | 112.6 |  |  |
|  | 3 | 112098 | 110.0 |  |  |
| 100% | 1 | 219952 | 107.9 | 109.3 | 1.1 |
|  | 2 | 224189 | 110.0 |  |  |
|  | 3 | 224130 | 110.0 |  |  |
| 150% | 1 | 341233 | 111.6 | 109.7 | 1.5 |
|  | 2 | 333045 | 109.0 |  |  |
|  | 3 | 331845 | 108.6 |  |  |

### Conclusion

- System suitability parameters were found to be within acceptance limits.

- EG and DEG were not detected in the unspiked (as such) samples, confirming specificity.

- At the LOQ level, recovery for both EG and DEG was within 70–130%, as expected.

- At 50%, 100%, and 150% spiking levels, the recovery ranged from 80–120%, under ICH Q2(R1) guidelines.

- %RSD for each level was less than 15%, confirming repeatability and reliability of the recovery procedure.

### 3.5 Solution Stability

The solution stability of the as-such sample and spiked sample solutions was evaluated by comparing the chromatographic responses of freshly prepared solutions with those stored in the autosampler tray over a period of 24 hours. Both Ethylene Glycol (EG) and Diethylene Glycol (DEG) were monitored for any significant change in content or chromatographic behaviour.

#### Preparation of Solutions

- Propylene Glycol Stock Solution
201.7 mg of Propylene Glycol was accurately weighed into a 100 mL volumetric flask. Approximately 70 mL of ethanol was added, the solution was mixed thoroughly, and the final volume was made up with ethanol.

- EG & DEG Stock Solution-A
100.9 mg of Ethylene Glycol and 100.7 mg of Diethylene Glycol were dissolved and diluted to 100 mL with ethanol.

- Standard Solutions A & B
5 mL of EG & DEG stock solution was diluted to 50 mL with ethanol, mixed well, and injected.

- System Suitability Solution
A mixture of 5 mL each of Propylene Glycol stock solution and EG & DEG stock solution was diluted to 50 mL with ethanol, mixed thoroughly, and injected.

- As Such Sample Solution
10.0138 g of sample was weighed into a 100 mL volumetric flask. Ethanol was added (approximately 70 mL), and the solution was sonicated for 10 minutes. The final volume was made up with ethanol, filtered using a 0.45 μm syringe filter, and injected.

- Spiked Sample Solution
10.0194 g of sample was weighed and spiked with 10 mL of standard stock solution in a 100 mL volumetric flask. Ethanol was added, and the solution was sonicated for 10 minutes. The final volume was adjusted with ethanol, filtered, and injected.

### Results

#### System Suitability

Table 17: System Suitability Parameters

| Parameter | Analyte | Observation | Acceptance Criteria |
| --- | --- | --- | --- |
| Similarity Factor (Standard A vs. B) | EG / DEG | 0.99 / 0.98 | 0.85 – 1.15 |
| Resolution between EG and Propylene Glycol | - | 12.82 | NLT 10 |
| % RSD of Peak Area (n = 6) | EG / DEG | 4.3 / 6.2 | NMT 15% |
| % RSD of Retention Time (n = 6) | EG / DEG | 0.03 / 0.03 | NMT 2% |
| Tailing Factor | EG / DEG | 1.7 / 1.9 | NMT 3.5 |

#### Stability Study Results

Table 18: Stability of Ethylene Glycol

| Time Point | Area | % Difference | Content (%) | % Difference |
| --- | --- | --- | --- | --- |
| Initial | 210958 | NA | 0.110 | NA |
| 2 hours | 214301 | -1.57 | 0.111 | -0.90 |
| 4 hours | 210759 | 0.09 | 0.110 | 0.00 |
| 8 hours | 211691 | -0.35 | 0.110 | 0.00 |
| 12 hours | 216628 | -2.65 | 0.113 | -2.69 |
| 18 hours | 226174 | -6.96 | 0.118 | -7.02 |
| 24 hours | 227025 | -7.34 | 0.118 | -7.02 |

Table 19: Stability of Diethylene Glycol

| Time Point | Area | % Difference | Content (%) | % Difference |
| --- | --- | --- | --- | --- |
| Initial | 218163 | NA | 0.112 | NA |
| 2 hours | 224838 | -3.01 | 0.115 | -2.64 |
| 4 hours | 212976 | 2.41 | 0.109 | 2.71 |
| 8 hours | 217302 | 0.40 | 0.111 | 0.90 |
| 12 hours | 214104 | 1.88 | 0.110 | 1.80 |
| 18 hours | 224205 | -2.73 | 0.115 | -2.64 |
| 24 hours | 226155 | -3.60 | 0.116 | -3.51 |

### Conclusion

- All system suitability parameters met the acceptance criteria, indicating a robust and reproducible method.

- The solution stability of both Ethylene Glycol and Diethylene Glycol was confirmed for a minimum of 24 hours.

- The observed differences in content and area responses were within ±15%, demonstrating solution stability as per analytical validation norms.

- No extraneous or degradant peaks were observed in any chromatogram.

- The % relative difference between initial and periodic interval results remained well below 60%, confirming chemical stability of the analytes in the matrix.

### Final Conclusion

Based on the results obtained from the method verification and validation study, the analytical method for the estimation of Ethylene Glycol and Diethylene Glycol in Naturals Cough Syrup is verified and found to be suitable for its intended purpose. All evaluated parameters, including specificity, precision, accuracy, linearity (if applicable), LOD, LOQ, robustness, and solution stability, comply with standard regulatory expectations (ICH Q2(R1)).

Acknowledgments

The authors sincerely thank the Delhi Analytical Research Laboratory for their invaluable technical support and for providing access to the instrumentation and infrastructure necessary for the successful completion of this work.

We extend our special gratitude to The University of Trans-Disciplinary Health Sciences and Technology for fostering a research-oriented environment and continuous academic encouragement.

A heartfelt thanks to my parents for their unwavering support, motivation, and understanding throughout the course of this study.

References

- ICH Q2(R1). (2005). Validation of Analytical Procedures: Text and Methodology. International Council for Harmonisation.

- USP <621> Chromatography. United States Pharmacopeia and National Formulary (USP-NF), USP 43–NF 38.

- FDA. (2015). Guidance for Industry: Analytical Procedures and Methods Validation for Drugs and Biologics. U.S. Department of Health and Human Services.

- European Medicines Agency. (2014). Guideline on Bioanalytical Method Validation.

- Snyder, L. R., Kirkland, J. J., & Dolan, J. W. (2010). Introduction to Modern Liquid Chromatography. John Wiley & Sons.

- Dong, M. W. (2006). Modern HPLC for Practicing Scientists. Wiley-Interscience.

- Kazakevich, Y. V., & Lobrutto, R. (2007). HPLC for Pharmaceutical Scientists. Wiley-Interscience.

- Swartz, M. E., & Krull, I. S. (2012). Analytical Method Development and Validation. CRC Press.

- Chou, S. S., & MacCrehan, W. A. (1995). Determination of ethylene glycol and diethylene glycol in pharmaceuticals using capillary gas chromatography. Journal of Chromatography A, 707(2), 317–322.

- EMEA. (2008). Diethylene Glycol in Contaminated Glycerin. European Medicines Agency report.

- Krishnaiah, Y. S. R., et al. (2013). Method development and validation for simultaneous determination of EG and DEG in cough syrup. Journal of Pharmaceutical and Biomedical Analysis, 85, 1–6.

- Dharuman, J., & Sinha, P. K. (2017). GC-FID method for determination of glycol contaminants in pharmaceutical preparations. Asian Journal of Chemistry, 29(12), 2579–2582.

- U.S. FDA. (2007). Public Health Advisory on DEG contamination in cough syrups. U.S. Food and Drug Administration.

- Jiang, Y., et al. (2010). Determination of multiple components in herbal cough syrup using GC-MS. Phytochemical Analysis, 21(6), 574–579.

- Mokhtari, V., et al. (2021). Herbal medicine quality control using chromatographic techniques. Evidence-Based Complementary and Alternative Medicine, 2021.

- WHO. (2005). Quality Control Methods for Herbal Materials. World Health Organization.

- EMA/HMPC. (2017). Guideline on quality of herbal medicinal products/traditional herbal medicinal products.

- Patel, R. M., & Patel, A. A. (2015). Challenges in analysis of polyherbal formulations. Journal of Pharmacy Research, 9(9), 541–545.

- Agency for Toxic Substances and Disease Registry (ATSDR). (2010). Toxicological Profile for Ethylene Glycol and Propylene Glycol.

- Barceloux, D. G., et al. (1999). Ethylene glycol poisoning: Mechanisms and management. Journal of Toxicology: Clinical Toxicology, 37(5), 581–609.

- Hantson, P., et al. (2000). Diethylene glycol poisoning: A case report. Human & Experimental Toxicology, 19(4), 231–234.

- Chan, Y. C., et al. (2004). Detection of DEG in adulterated cough syrups. Journal of Forensic Sciences, 49(5), 1058–1061.

- Samanidou, V. F., et al. (2004). Development and validation of an HPLC method for determination of preservatives and additives. Journal of Chromatographic Science, 42(6), 295–301.

- Nogueira, J. A., et al. (2002). Validation of GC-MS methods for trace analysis in complex matrices. Talanta, 58(5), 1065–1072.

- Rao, R. N., et al. (2006). RP-HPLC method validation for pharmaceutical compounds in complex matrices. Journal of Pharmaceutical and Biomedical Analysis, 40(3), 659–667.

- Grob, R. L., & Barry, E. F. (2004). Modern Practice of Gas Chromatography. Wiley-Interscience.

- Marriott, P. J., et al. (2001). Comprehensive two-dimensional gas chromatography. Journal of Chromatography A, 936(1–2), 1–22.

- Dettmer-Wilde, K., & Engewald, W. (2014). Practical Gas Chromatography: A Comprehensive Reference. Springer.

- WHO Alert (2022). Medical Product Alert N°6/2022: Contaminated Cough Syrups. Geneva: World Health Organization.

- GACVS (Global Advisory Committee on Vaccine Safety). (2023). Safety of medicinal syrups and glycol contamination.

- CDSCO (India). (2023). Regulatory Circular on DEG Contamination in Pharmaceuticals.

- Health Canada. (2022). Advisory on DEG contamination in children's cough formulations.

- Kumar, A., et al. (2015). Quality control of liquid oral formulations using chromatographic tools. Indian Journal of Pharmaceutical Sciences, 77(4), 475–481.

- Raj, K. G., et al. (2019). Validation of analytical methods in herbal matrix using QbD approach. International Journal of Green Pharmacy, 13(3), 213–219.

- Nema, R. K., & Kumar, A. (2008). Quality assessment of polyherbal syrups: An overview. Pharmacognosy Reviews, 2(4), 62–68.

- Sharma, S., et al. (2020). Current status of herbal drug analysis and validation. Asian Journal of Pharmaceutical Analysis, 10(2), 102–110.

- Satheeshkumar, N., et al. (2014). Challenges and strategies in herbal drug quality control. Journal of Natural Remedies, 14(2), 85–94.

- Kaur, H., et al. (2021). Analytical method development for herbal APIs using chromatography. Journal of AOAC International, 104(1), 182–191.

- Chawla, G., & Singh, S. (2009). Impurity profiling: A regulatory perspective. Pharmaceutical Technology Europe, 21(11), 40–48.